Object:  		STANDARD OPERATING PROCEDURE

Document title:  	DNA extraction from intestinal samples

Version: 		0.1

Department:		IKBM

Author:		Inga Leena Angell/Technician

Date: 			10.05.16

Reviewed by:		Knut Rudi

1 Purpose

The purpose of the present Standard Operating Procedure (SOP) is to describe how the DNA is isolated from intestinal samples by bead-beating, and a LGC mag midi DNA extraction kit (LCG genomics) in a KingFisher Flex robot (Thermo Scientific).

2 Procedure

2.1 	Receiving stool samples

Stool samples should be sent on dry ice and kept frozen at all times.

2.2	Buffers, reagents and solutions

Table 2.1 Overview of buffers, reagents and solutions for DNA isolation of intestinal content.

| Buffer/reagent/kit | Product no. | Supplier |
| --- | --- | --- |
| Mag midi kit | 40402/40420 | LCG genomics |
| Acid washed beads (<106 µm) |  | Sigma-Aldrich |
| STAR buffer |  | Roche |

2.3	Equipment

| Equipment | Vendor/Supplier |
| --- | --- |
| KingFisher Flex robot | Thermo Scientific |
| KingFisher Flex KF plates | Thermo Scientific |
| KingFisher Flex tip combs | Thermo Scientific |
| MagNA Lyser | Roche |
| Table top centrifuge | General supplier |
| Qubit fluorometer | Invitrogen |
| Sterile screw cap tubes, 2ml | Sarstedt |
| Standard equipment (vortex, micropipettes etc.) | General supplier |

2.4	Thawing fecal samples

The samples must be thawed on ice.

Aliquot 0.25 g acid washed glass beads (Sigma-Aldrich, < 106 µm) into 2.0 ml into sample tubes with screw cap and silica ring in the lid. Add 200 ul S.T.A.R buffer into each tube (the v/v ratio should be 2/1 STAR/sample). Keep the samples on ice at all times.

2.5	DNA isolation

When all the stool samples are aliquoted into the tubes, you should immediately continue with the procedure for bacterial DNA isolation from the stool samples.

Procedure for bacterial DNA isolation from stool samples.

- Mechanical lysis

1.1 Process twice in MagNaLyser at 6500 rpm for 20 sec with 1 min rest between runs. Keep samples cold during the rest.

1.2 Centrifuge the tubes at 13000 rpm for 5 min.

- Proteinase treatment

2.1 Transfer 50 µl supernatant from the tubes to KingFisher 96 well plate

2.2 Add 50 µl lysis buffer and 5 µl Proteinase – mix well by pipetting

2.3 Place the plate into KingFisher Flex robot. Use ‘ProteinaseLGC’ procedure to incubate the   samples at 55 ⁰C for 10 min.

2.4 Store on ice until further use.

3. DNA extraction

Add reagents only to the wells that will actually contain sample.

Prepare KingFisher plates according to table:

| Plate Name | Plate type | Content | Volume |
| --- | --- | --- | --- |
| Sample Plate | KingFisher 96 | Mag particles Sample (from step 4.4) Ethanol (96 %) | 16 µl 105 µl 50 µl |
| Wash Plate 1 | KingFisher 96 | BLM 1 | 170 µl |
| Wash Plate 2 | KingFisher 96 | BLM 2 | 175 µl |
| Wash Plate 3 | KingFisher 96 | BLM 2 | 175 µl |
| Elution Plate | KingFisher 96 | Elution buffer | 50 µl |
| Tip Plate | KingFisher 96 | Tip comb |  |

Start procedure “MagMiniLGC”.

Place the six plates in the instrument according to instructions on the display.

When the run has ended, transfer the samples into a 96-well storage plate.

2.6	Quantification of genomic DNA

Measure by Qubit fluorometer. Use 2 ul for each measurement.

2.7	Storage of DNA

After DNA isolation the sample should be stored with appropriate labelling. For short term-storage (up to 3 days), store at +4 oC. For long term storage (more than 3 days), store at -20 oC.

3	References

Magnalyzer:

https://www.roche-applied-science.com/sis/automated/lyser/index.jsp

Mag midi kit:

http://www.lgcgroup.com/LGCGroup/media/PDFs/Products/Extraction/Kits%20inserts/mag-midi.pdf

KingFisher Flex robot:

https://tools.thermofisher.com/content/sfs/manuals/D01475~.pdf

Qubit fluorometer:

https://www.thermofisher.com/content/dam/LifeTech/migration/en/filelibrary/cell-tissue-analysis/qubit-all-file-types.par.27078.file.dat/qubit-usermanual.pdf

Qubit assay protocol:

https://www.thermofisher.com/content/dam/LifeTech/migration/en/filelibrary/cell-tissue-analysis/qubit-all-file-types.par.90790.file.dat/qubit1.0-protocol.pdf